FZ-1/5509/AJ/21										Annex No. 1

Technical requirements:

The subject of the preliminary inquiry is the purchase of dosemeters, holders for dosemeters and photomulitiplier tube.

The offered parts must be brand new, adapted to use with owned by the Awarding Entity Panasonic dosemeters system, consisting of an automatic reader UD-7900M and automatic equipment UD-794D for irradiating dosemeters and must to meet the following requirements:

| Description | Quantity pcs. |
| --- | --- |
| 4 elements UD–802AT–D TLD dosemeters with barcode, compatible: with holders for dosemeters UD–874ATM TLD and with Panasonic reader UD-7900M owned by the Awarding Entity. | 2 000 |
| Plastic holders UD–874ATM TLD for dosemeters with separate metal grip – clip, compatible with Panasonic 4 elements UD–802AT–D TLD dosemeters. | 2 300 |
| Photomulitiplier tube R464-05 Hamamatsu, compatible with Panasonic automatic reader UD-7900M, owned by the Awarding Entity. Application of the photomultiplier tube: for counting photons. Specification: minimum effective area of the photocathode 5 x 8 mm; supply voltage between the anode and the cathode 1500 V; average anode current 0,01 mA; rise time of the anode pulse 13 ns; electron transit time 70 ns. | 1 |